5135A Universal Digital Panel Meter

I  Product Highlights 

This Meter is a 3 ½ Digits Panel Meter, is great for measurements that require high precision and stability. It is widely used to measure and display of voltage, current as well as other critical parameters in variety of electrical equipments and instruments.  Surface Mount Technology (SMT) has been adopted for meter’s PCB mass fabrication, which greatly increases the product quality.

※1.  Can be used as Voltmeter, Amp meter, or general readout display.

※2.  Accuracy adjustment.

※3.  Decimal point position adjustable.

※4.  Digital filter at input can be turned ON or OFF. 
※5.  Hold current reading is available.
 Capable of measuring its own supply voltage source.  

II Specifications and maim parameters

[image: image1.jpg]FAGA.

%2
Calibration

ColdfusionX




Operating Voltage:  DC5V +/-10%

Operating Current:  ≤ 80(mA)

Display Range:  0 to ± 1999

Display LED Height:  18.6(mm) or 9/16(inch)

Displayed when over limit:  “EEEE” or “-EEE”

Sampling period:  0.3sec

Error:  ± 0.5%

Operation Environment: 

≤ 85% Relative Humidity

            0C ~ 50C, (32F ~122F), Ambient Temperature

Meter External Dimensions:  79x43x20(mm)  or  3.11x1.69x0.79(inch)

Opening for Meter installation:  76x39.5(mm)  or  2.99x1.56(inch) 

III Operation Instruction

a) Accuracy Adjustment: measure a know voltage and turn the knob on the blue variable resistor to the know value.  Adjustable range would be around +/-100. 

b) Display Memory: Short the two points where marked “M” to activate memory.

 c) Digital Filter: Short the two points where marked “C” to activate digital function, which can filter out noises and input signal fluctuation.   

d) Decimal Point setting: By using D1,D2 and D3.

When nothing shorted, no decimal point is displayed.

When D1 is shorted, decimal is between two right most digits

When D2 is shorted, decimal is between two center digits

When D3 is shorted, decimal is between two left most digits. 

e) Expansion of measurement range: 
Obtain right measurement range by changing resistors according to the following table.

	 
	Input Signal
	Display Range
	Required Values

	
	 
	 
	R1
	R2
	R3

	As Voltage Meter
	0 ~ ±20V
	0 ~ ±19.99
	1M
	 
	10K

	
	0 ~ ±200V
	0 ~ ±199.9
	1M
	 
	1K

	 

As Current

Meter
	0 ~ ±2mA
	0 ~ ±1.999
	100K
	100Ω
	 

	
	0 ~ ±20mA
	0 ~ ±19.99
	100K
	10Ω
	 

	
	0 ~ ±200mA
	0 ~ ±199.9
	100K
	1Ω
	 

	
	0 ~ ±2A
	0 ~ ±1.999
	100K
	0.1Ω (2W)
	 

	 

 

 

As Generic

Meter
	0 ~ ±5V
	5V Full Scale
 0 ~ ±1999
	1M
	 
	[100(Display Value) ÷(5K-Display Value)]KΩ

	
	0 ~ ±10V
	10V Full Scale 0 ~ ±1999
	1M
	 
	[100(Display Value) ÷(10K-Display Value)]KΩ

	
	0 ~ ±10mA
	10mA Full Scale 0 ~ ±1999
	100K
	(F.S.÷100)Ω
	 


Resistor requirements:  
Resistors with 5.0% error rating, some adjustment of the variable resistance might be needed to obtain accuracy.  Resistors with 1.0% error rating, can be directly used without adjustment.  
[image: image2.jpg]FPanel Meter 5135
0 ~+50. 0/75mV

Fanel Meter 5135
0~+199. 9/200V

Fanel Meter 5139

Fanel Meter 5135
0 ~+50. 0/75mV

0~+199. 9/200V

Tz 12
48.0V e | 48.0V 5
50. 0A Output 50. 0A Output
Stable DG48V —Lm ov Stable DC48V ov
Power Supply| o Power Supplyl  ghnt
50A/75mV e B - It 50A/75mV = &
Power Power Power
Supply Supply Supply
5V,/200mA 5V/200mA 5V/200mA
Figure 1

Figure 2




